SAMPLE POLICY:  ETHOS SOP

Policy Name: Ethos® Standard Operating Procedure

Last Review Date: 04/01/2026                                         Date of Origin: 01/01/2026

Scope: Applies to all hospitals, clinics, and ambulatory areas in the healthcare system.

Policy: The healthcare system is committed to safe and effective reprocessing of endocavity and surface probes through standardized & automated cleaning and high-level disinfection (HLD).

Purpose: To promote compliance, patient safety, and infection prevention with accepted guidelines for the reprocessing of endocavity and surface probes between patient use.

Key Words:

Point of Use Treatment: the prompt, initial treatment steps and/or measures to prevent drying of soil and contaminants in and on the device, as well as to prepare the device for further processing

Note to entry: Previously referred to as pre-cleaning, the adoption of the term "point of use treatment" is for greater clarity as to the intention of this step. Additional measures could include disconnecting accessories, etc.

Point-of-use: occurring at the bedside or where the procedure was performed

Cleaning: removal of contamination from an item to the extent necessary for further processing or for the intended use

Personal Protective Equipment (PPE): according to the Code of Federal Regulations, “specialized clothing or equipment worn by an employee for protection against hazardous materials, blood-borne pathogens, or other potential infectious materials. General work clothes (e.g., uniforms, pants, shirts or blouses) not intended to function as protection against a hazard are not considered to be personal protective equipment” (29 CFR 1910.1030)

Low-level disinfection (LLD): lethal process utilizing an agent that kills vegetative forms of bacteria, some fungi, and lipid viruses

High-Level Disinfection (HLD): a process that kills all microbial organisms, but not necessarily large numbers of bacterial spores

Enzymatic: specialized solutions used for cleaning, typically made up of proteases, amylases, and lipases—designed to break down organic soils, such as blood, tissue, fats, and biofilms.

Biohazard: any used instrument, tool, or consumable item contaminated with blood, bodily fluids, or other potentially infectious materials (OPIM) that poses a risk of transmitting pathogens (such as HIV, hepatitis, or bacteria) to humans or the environment

Don: securely putting on protective gear (e.g., gown, mask, goggles, gloves) before potential exposure

Doff: the systematic, safe removal of PPE, often in a specific order to avoid self-contamination, typically from dirtiest to cleanest.

HEPA filtration: a high-efficiency mechanical air filtration system that removes at least 99.97% of airborne particles, bacteria, and potential contaminants with a size of 0.3 microns (µm) or larger from the air inside a storage cabinet

Procedure:

Pre-Cleaning:

- Don appropriate PPE per facility policy.

- Perform pre-cleaning at the point of use, per the IFU, before blood, gel, body fluids, or bioburden have a chance to dry. Remove the sterile sheath/condom during the pre-cleaning process.

Soiled Transport:

- To prevent cross-contamination and probe damage, place the probe in a biohazard Probe ValetTM bag to separate the parts that need high-level disinfection (HLD) from the cord and electrical pack that need low-level disinfection (LLD). Seal the top of the bag.

- Doff and discard PPE. Cleanse hands.

Cleaning & High-Level Disinfection

- Don appropriate PPE, per facility policy.

- Place the endocavity/surface probe inside the Ethos chamber. Close the chamber door.

- Scan the operator’s name and the probe serial number.

- Insert the AquaCide® cleaner-disinfectant bottle into the Ethos equipment.

- Start Cleaning & HLD cycle.

- LLD cord and electrical pack.

- At completion of the cycle, record PASS/FAIL results in the logbook.

  - Note the next vapor filter and 5nm water filter change dates, printed out on cycle receipt.

- Doff PPE. Cleanse hands.

Drying & Storage:

- Don clean gloves.

- Using a QwikDry® cloth (single-use, single-packaged, low-linting drying cloth), remove the endocavity/surface probe from the Ethos machine and dry the entire length of the probe handle and shaft.

- Hang the endocavity/surface probe vertically in a storage cabinet designated specifically for that type of probe, ensuring the HLD and LLD components do not come in contact with each other.

- Doff gloves. Cleanse hands.

- A storage cabinet should have HEPA-filtered air circulation and a locking mechanism on the door.

- Use a Hang Time TagTM on each probe to indicate when the probe's storage time has expired, and reprocessing is required.

- Reprocess each probe once 7 days of storage have elapsed.

- Cabinets used for probe storage should be cleaned and disinfected with hospital-approved disinfectant when visibly soiled, and weekly.

Training & Competency:

- Personnel involved in probe processing using HLD must be provided education, training, and complete competency verification activities related to their duties at a minimum, at the following times:

  - Upon initial hire

  - Annually

  - When new equipment, probes, or supplies are brought into the department

  - At any time designated by the department leader

- Processing activities should be closely supervised until competency is verified and documented for each processing task, from cleaning to probe storage.

References:

Association for the Advancement of Medical Instrumentation. (2017). ANSI/AAMI TIR99:2024 – Processing of dilators, transesophageal, and ultrasound probes in healthcare facilities.  Arlington, VA.

CSIB.0051.260415