SAMPLE POLICY: TEECLEAN SOP

Policy Name: TEEClean® Standard Operating Procedure

Last Review Date: 04/01/2026			Date of Origin: 01/01/2026

Scope: Applies to all hospitals, clinics, and ambulatory areas in the healthcare system.

Policy: The healthcare system is committed to safe and effective reprocessing of transesophageal echocardiogram (TEE) probes through standardized and automated cleaning and high-level disinfection (HLD).

Purpose: To promote compliance, patient safety, and infection prevention with accepted guidelines for the reprocessing of TEE probes between patient use.

Key Words:

Point of Use Treatment: the prompt, initial treatment steps and/or measures to prevent drying of soil and contaminants in and on the device as well as to prepare the device for further processing

Note to entry: Previously referred to as pre-cleaning; the adoption of the term "point of use treatment" is for greater clarity as to the intention of this step. Additional measures could include disconnecting accessories, etc.

Point-of-use: occurring at the bedside or where the procedure was performed

Cleaning: removal of contamination from an item to the extent necessary for further processing or for the intended use

Personal Protective Equipment (PPE): according to the Code of Federal Regulations, “specialized clothing or equipment worn by an employee for protection against hazardous materials, blood-borne pathogens, or other potential infectious materials. General work clothes (e.g., uniforms, pants, shirts or blouses) not intended to function as protection against a hazard are not considered to be personal protective equipment” (29 CFR 1910.1030)

Low-level disinfection (LLD): lethal process utilizing an agent that kills vegetative forms of bacteria, some fungi, and lipid viruses

High-Level Disinfection (HLD): process that kills all microbial organisms but not necessarily large numbers of bacterial spores

Enzymatic: specialized solutions used for cleaning, typically made up of proteases, amylases, and lipases — designed to break down organic soils, such as blood, tissue, fats, and biofilms

Biohazard: any used instrument, tool, or consumable item contaminated with blood, bodily fluids, or other potentially infectious materials (OPIM) that poses a risk of transmitting pathogens (such as HIV, hepatitis, or bacteria) to humans or the environment

Don: securely putting on protective gear (e.g., gown, mask, goggles, gloves) before potential exposure

Doff: the systematic, safe removal of PPE, often in a specific order to avoid self-contamination, typically from dirtiest to cleanest.

HEPA filtration: a, high-efficiency mechanical air filtration system that removes at least 99.97% of airborne particles, bacteria, and potential contaminants with a size of 0.3 microns (µm) or larger from the air inside a storage cabinet

Procedure:

Pre-Cleaning:

- Don appropriate PPE per facility policy.

- Perform pre-cleaning at point-of-use, before blood, body fluid or bioburden has an opportunity to dry.

- Use the pre-packaged TEEZyme® enzymatic sponge to wipe up and down the shaft of the TEE probe, avoiding twisting motion, three times.

Soiled Transport:

- In order to transport in a manner that prevents cross-contamination and probe damage, cover the probe with a red PullUp™ probe cover (to indicate a dirty/biohazardous probe), place a TDefender™ on the tip for protection and place in the TPorter® procedure case for transport to the reprocessing area.

- Doff and discard PPE.  Cleanse hands.

Cleaning & High-Level Disinfection

- Don appropriate PPE, per facility policy.

- Remove soiled TEE probe from TPorter procedure case. Remove biohazard PullUp from TEE Probe.

- Place TEE probe inside the TEEClean chamber. Close chamber door. Follow prompts on the LCD screen of TEEClean.

- Scan the operator’s name and the probe serial number.

- Perform Electrical Leak Testing, record results in logbook.

- Insert either TD-5® or TD-8® disinfectant bottle into TEEClean equipment.

- Start Cleaning and HLD cycle.

- LLD cord and electrical pack.

- At completion of the cycle, record PASS/FAIL results in logbook.

  - Note the next vapor filter and 5nm water filter change dates, printed out on cycle receipt.

- Doff PPE. Cleanse hands.

Drying & Storage:

- Don clean gloves.

- Using a QwikDry® wipe (single use, single packaged, low-linting drying cloth), remove the TEE probe from the TEEClean machine and dry the entire length of the probe shaft.

- Hang TEE probe vertically in a HEPA filtered storage cabinet designated specifically for TEE probes, ensuring the HLD and LLD components do not come in contact with each other.

- Doff gloves. Cleanse hands.

- Storage cabinet should have HEPA circulated air and a locking mechanism on the door.

- Use a Hang Time TagTM on each probe to designate when the probe storage time has expired (at 7 days) and reprocessing is required again.

- Storage cabinets used for TEE probes should be cleaned and disinfected with hospital approved disinfectants when visibly soiled and on a weekly basis.

Training & Competency:

- Personnel involved in TEE probe processing using HLD must be provided education, training and complete competency verification activities related to their duties at minimum, at the following times:

  - Upon initial hire

  - Annually

  - When new equipment, probes, or supplies are brought into the department

  - At any time designated by the department leader

- Processing activities should be closely supervised until competency is verified and documented for each processing task, from cleaning to storage of the TEE probe.

References:

CSIB.0050.260410

Association for the Advancement of Medical Instrumentation. (2017). ANSI/AAMI TIR99:2024 – Processing of dilators, transesophageal and ultrasound probes in healthcare facilities.  Arlington, VA.